The 8051 Data Collection Unit

(0) Introduction

This reference details the implementation and operation of the data

collection software included in this package, and its required hardware

configuration.

(1) Microprocessor-Based Data Collection

The original purpose of this software was to carry out a set of fluid flow

tests on fluid meters in parellel.  Two of these units could be controlled from

a PC over a RS-485 serial link.  Each unit could test up to 6 meters at a time.

Each meter has inside it a magnet which rotated in response to the fluid

flowing through it.  At a given flow rate, this magnet will rotate a constant

amount for each unit of fluid volume passing through the meter.  This constant

will be approximately, but not exactly, the same for different flow rates and is

an inherent characteristics of the meter's construction.

The magnet internal to the meter is divided into a fixed number of sectors

of alternating magnetic field.  As the magnet rotates, it can be used to

generate an square electrical waveform which is ideally suited for measurement

by a microprocessor.

The microprocessor-based data collection unit will mark off a set number

of square waves (corresponding to one full rotation of the magnet), measure

its width, and collect the information for successive measurements.  It will

use this information to determine the number of complete rotations of the

meter's magnet and to estimate the number of fractional rotations, thus

obtaining a highly accurate measurement of the meter's behavior.

For future reference, the term NUTATION is also used to mean one full

rotation of the meter's internal magnet.

The system is set up in such a way that each unit will receive a signal at

the start of a test and a signal at the end of the test.  The function of these

signals is to delimit a test and mark off a fixed volume of fluid.

The remaining part of this reference will refer to this specific

configuration, but it should be noted that this software is completely generic.

It can be used to measure and perform statistical processing on any set of up

to 6 square wave inputs using a 7th falling edge input to mark the start and end

of a test.

(a) Hardware Interface

Data collection is done using an Intel 8051FA microprocessor running

dedicated stand-alone software.  Each microprocessor is connected to 6 pulse

inputs, to a sensor signal input, and to a RS-485 serial communications

connection.  The pulse inputs enable the microprocessor to perform measurements

on each of the 6 sources it is monitoring, as explained above.  The sensor

input is responsible for signalling to microprocessor when to start and stop

the measurement and counting.  Finally, the communications link enables the

micrprocessor to communicate with the master PC so that its data can be

transferred to the central controller.

The inputs and outputs are set up as listed below:

Function                          Location

Pulse input 0                 T2EX: Timer 2 External

Pulse input 1                 CEX0: Counter External 0

Pulse input 2                 CEX1: Counter External 1

Pulse input 3                 CEX2: Counter External 2

Pulse input 4                 CEX3: Counter External 3

Pulse input 5                 CEX4: Counter External 4

Sensor input                  INT0: External Interrupt 0

Serial Communications         RxD: Receive input

TxD: Transmit output

T0: RS-485 transmit enable

(b) Internal Software Architecture

This is a multi-tasking software architecture.  Several independent tasks

will be running simultaneously on the one 8051FA chip.  They communicate with

each other and with the outside world using the 8051 multitasking library.

Because of the large number of processes and the limit on stack space, several

compromises were made when integrating this library with the rest of the data

collection software.

The microprocessor is responsible for monitoring 6 counters, and performing

statistical operations for each counter concurrently as it is collecting the

data.  A processing loop is set up for each counter and is synchronized by the

counter input.

A process is set up to monitor a test, and is synchronized to the sensor

input.  When the first signal is receives, this process will Spawn the 6

counter processes (and also 2 timer processes).  When it receives the second

signal, it will disable the counter inputs (and clocks) and thus stop the

processes it spawned.

Finally, the microprocessor will monitor the serial communications link,

waiting for commands from the master PC.  These commands serve three purposes:

(a) the enable/disable the master test process, (b) the signal the

microprocessor to transfer its data over the serial communications link, and

(c) for diagnostic purposes to test the microprocessor's status and proper

functioning.

(c) Data Collection Method

The data collection method used measures the times for full meter nutations

during a test between the two sensor signals

Suppose these events take place, as illustrated in the time-line below:

*---------------------------------------------------------------------------*

| N = Meter Nutation                                                        |

| S = Sensor Signal                                                         |

|                                                                           |

|                  S                                         S              |

|        N   N   N     N      N      N      N      N     N     N    N       |

|                  *-----------------------------------------*              |

|        |   |   | | F |  w0  |  w1  |  w2  |  w3  |  w4 | L | |    |       |

|                  *-----------------------------------------*              |

|                  ^   ^                                 ^   ^              |

|                  |   |   <--- N full nutations --->    |   |              |

|                  S   B                                 C   T              |

|                                                                           |

|                  Events: S = First Sensor Time,                           |

|                          B = First Meter Nutation Time,                   |

|                          C = Last Meter Nutation Time,                    |

|                          T = Last Sensor Time                             |

|                                                                           |

|   First Partial Pulse Width: F = B - S                                    |

|   Last Partial Pulse Width:  L = T - C                                    |

|   Average Pulse Width:      <w> = (C - B)/N (= (C - B)/5 in this example) |

|                                                                           |

|   Estimated First Partial: F/<w>                                          |

|   Estimated Last Partial:  L/<w>                                          |

|   Estimated Total Nutations:      N + F/<w> + L/<w>                       |

|     which actually comes out to:  (T - S)/<w> =  N (T - S) / (C - B)      |

*---------------------------------------------------------------------------*

Then the tester will only be actively counting BETWEEN events S and T.  The

data collection method will calculate the nutation widths w0, w1, w2, w3, w4

(thus counting 5 full nutations in this example); and F and L -- the first and

last partial nutations.  It will then ESTIMATE what fraction of full nutations

F and L represent by dividing F and L respectively by the average value of the

w's.

To ensure that a steady flow condition is valid, we also measure the VARIANCE

of the pulse widths, which is defined here as:

Variance = Standard-Deviation of w's / Average of w's  x 100%

Typical values in the actual test system this software was used for could vary

between 2% and 3% for high flow rates, down to 0.5% for low flow rates.  A high

variance (say, above 5%) invalidates the extrapolation method we use, which is

premised on a constant flow rate.  High variances can happen for two reasons:

(1) Fluid flow is not in a steady state at the test boundaries.

This can occur if the underlying system is still in a transient state

as a test starts.

(2) The unit undergoing test itself is not operating reliably.

(d) Counter Processing

To be able to deal with 6 counters with one microprocessor requires some

special measures be taken.  On average, for our actual test conditions (20

nutations per second received from each of the 6 inputs) the 8051FA will spend

30% of its time during the high flow test doing actual statistical calcultions.

However, the demand on the processor can vary between 5% and 200% depending on

when the 6 inputs delimit full nutations relative to one another.  The maximum

load will occur when the majority of the 6 inputs have full nutations close to

one another in time.

What we do is the following.  When an input has signalled one full nutation,

the time of this event is saved in a protected location.  A completely

independent process, called the Scheduler, started at the beginning of the test,

will periodically remove one AND ONLY ONE of these saved items, passing it on

to the corresponding statistical processing routine.  This simple, but

effective, strategy will keep the processor active at about a constant 30% rate

(for the test conditions indicated above).

(e) Communications

The details of serial commuications will be described in more detail in

the section dealing with the MicroNet.  The underlying protocol used for this

system is: 9 data bits, 1 stop bit, no parity, and 9600 baud (9600,N,9,1).

The 9th data bit is used for networking purposes to distinguish commands from

data.  The centrally located PC will always send commands with the 9th bit high,

and the microprocesors will always send data, with the 9th bit low.  This

protocol will help protect the network from inadvertent access by two or more

processors at once.

(f) Command Reference

As a slave unit to the central PC, each microprocessor will respond to a

set of commands.  These commands are described in brief below and will be

detailed in the section on the MicroNet.

STATUS: A diagnostic command that queries the current state of the data

collection unit.  It will respond to the PC that it is in one of

the following modes:

ACTIVE: In active command mode, ready to start a new test.

WAITING: In test mode, waiting on the first sensor.

TESTING: In test mode, collecting data, waiting on the second sensor.

TEST:  A command telling the microprocessor to pass into the waiting mode.

It is ignored unless the microprocessor is active.

ABORT: A command that forces the processor back into the active mode and

cancels the current test.

DUMP:  A diagnostic command used by the PC to obtain the detailed set of

nutation widths collected by the processor in the most recent test.

STATS: A command used by the PC to obtain the statistical results of the

most recent test.

During the normal course of system operation, only the TEST and STATS commands

will ever be used.  The ABORT is used on exceptional occasions such as when a

communication failure takes place, or when a test is aborted by the PC.  The

other two commands are only used by diagnostic programs.

(2) MicroNet

MicroNet is a relatively simple multi-processor network custom-designed

specifically for use with this test system.  It has a multi-drop topology

in which one master controller (a PC) communicates with several slave

processors (8051FA's).  The network has been designed in several layers,

as is standard practice: a Physical Layer, a Protocol Layer, and a Command

Layer.

This section has all the details needed for anyone who wants to write client

software at a PC (or elsewhere) to communicate with the microprocessors.

(a) The Physical Layer

This layer is concerned primarily with the actual details of the connections

between the processors.  An RS-485 serial communications link is being used,

which is similar to the RS-422 link, except it allows several processors to be

on the line, instead of just two.

The actual circuit is a loop terminated by a resistor.  Each processor on the

loop receives input from the network by sampling inputs from opposite ends of

the loop across the resistor.

At any given time, up to one processor can take control of the loop, passing

into a transmission state.  All other processors must be in a reception state

at that time for proper signalling to take place.  For this purpose, each

processor has a transmit/receive control output which allows it to connect or

disconnect its transmitter.

The microprocessor units use the T0 pin, mentioned briefly in the previous

section for this purpose.  The master PC must likewise use its own control

output for the same function.  Normally, an RS-485 serial communications card

will need to be installed in the PC.

(b) The Protocol Layer

The primary purposes of the protocol layer are to:

(1) Protect the network against multiple access by two or more

transmitters.

(2) Ensure proper routing of transmissions.

(3) Ensure reliable transmission.

This is actually done at two levels: one closely tied to the physical layer

and the other implemented completely in software.

(i) Low-level, physical, protocol

For transmission, we use the protocol 9600, N, 9, 1: 9600 baud, no parity,

9 data bits, and one stop bit.  Error-free communication has been proven at baud

rates up to 57600 with the units all within 2 meters of each other.

The first 8 data bits are used to transfer information or commands.  The 9th

bit is used for protocol purposes in the following way:

* All transmission to a microprocessors occurs with the 9th bit set to 1.

* All transmission from a microprocessor occurs with the 9th bit set to 0.

* All microprocessors ignore transmissions where the 9th bit is 0.

These three conditions make it logically impossible for microprocessors to

communicate with one another, when the underlying Physical Layer is functioning

reliably.  They can only communicate with the master PC.  This is one of the

key conditions that assures the network's protection against simultaneous

access.  Complete assurance is provided by the software portion of the protocol.

(ii) High-level, software, protocol: Addressing

Since there are only two micrprocessor units on the network, routing is

greatly simplified.  Call these units A, and B.  All commands to the

microprocessors will have the following format:

1   B A   C C C   M M M

The high-order bit is always set to 1, as explained above.  The next two

bits (A, and B) are ADDRESS bits.  They have the following definitions:

A: Indicates unit A as a recipient of the transmission

when set.

B: Indicates unit B as a recipient when set.

All tranmissions to unit A will therefore have the pattern: 1 x1 xxx xxx; those

to unit B, the pattern 1 1x xxx xxx; and those directed to both units, the

pattern: 1 11 xxx xxx.

The bit pattern 1 00 xxx xxx is universally ignored by the microprocessors.

It has been used allow communications with other microprocessors that may be

installed in the communications line for diagnostic purposes (for example,

a diagnostic microprocessor that simulates sensor signals).

The next 3 bits (C C C) are the COMMAND FIELD, and the last 3 (M M M) the

COUNTER/SUBCOMMAND field.  They are used by the Command Layer to implement the

basic commands described in section 1f.

(iii) Data Transfer Formats

To provide an added check against transmission errors, two data transmission

formats are defined: one for short transmissions and the other for long

transmissions.  When blocks of data are to be transferred from a microprocessor

one of these formats will be used.

For short transmissions (in response to the STATS command defined in section

1f), the following bytes are transferred:

'#': Indicates start of a short transmission

SIZE: Indicates the amount of data being transferred.

This can range from 1 to 100 (hexadecimal).  The

latter value is indicated when SIZE is set to 0.

<Data>: A number of bytes indicated by SIZE

Check Sum: The sum of all the Data bytes, modulo 100 (hex).

Long transmissions (used for the DUMP command) are formatted the same,

except that the header is ':', instead of '#'.  Also, with long transmissions,

the master PC is expected to send back one of the following responses:

ACCEPT:  The bit sequence: 1 xx 011 000

REJECT:  The bit sequence: 1 xx 011 011

STOP:    The bit sequence: 1 xx 011 111

Any other bit sequence received by the microprocessor is interpreted as a STOP,

for the sake of network security.

If the microprocessor receives a REJECT, it will resend the data it just got

through sending.  If it receives an ACCEPT, it will send the next block of data.

If it receives a STOP, it will abort the long transmission.  A pre-determined

number of consecutive REJECT's (3) is interpreted as a STOP.

After receiving an ACCEPT on the last block of data transmitted, the micro

will signal the end of transmission by sending a period: '.'.  This indicates

the successful completion of a long transmission.

This protocol is a simplification of the XMODEM protocol.

(c) The Command Layer

To implement the commands described in section 1f, the following bit

patterns are used:

Commands to the microprocessor:

STATUS              1  x x  0 1 0  0 0 0

TEST                1  x x  0 1 1  0 0 0

ABORT               1  x x  0 1 1  1 1 1

DUMP from Input m   1  x x  0 0 1  m m m

STATS from Input m  1  x x  0 0 0  m m m

with the following key

mmm          Input

000            0

001            1

010            2

011            3

100            4

101            5

A summary of these commands is given below:

STATUS: Queries the current state of the data collection unit.  The

microprocessor will respond sending one of the following characters:

'0' = ACTIVE: In active command mode, ready to start a new test.

'1' = WAITING: In test mode, waiting on the first sensor.

'2' = TESTING: In test mode, collecting data, waiting on the second

sensor.

TEST:  Tells the microprocessor to pass into the waiting mode.

Ignored unless the micro is in the ACTIVE mode.

ABORT: Forces the processor back into the active mode.  Also used to abort

long data transfers.

DUMP:  Initiates a transfer by the microprocessor to the master PC of the

nutation widths for the indicated counter of the most recent test, using

the long transfer format.  Each nutation width is send as an unsigned

4-byte integer with the lowest order byte being sent first.

STATS: Initiates a transfer to the PC of the statistical data of the most

recent test.  The following data is packaged into the short transmission

format and is then sent:

Bytes      Name     Meaning

0       State     Bit 6: set if there was no test since the micro was

reset, or if the last test was aborted.

Bits 0-5: Set if the respective input 0-5 never

received any input from a meter.

1-2      Cycles    The number of complete nutations counted in the test

3-6      TIME      The total test time

7-10     First     The time that the first nutation occured

11-14     Last      The time that the last nutation occured

15-22     Square    The sum of the squares of all the nutation widths.

Corresponding to the figure shown in section 1c, the last 5 items are,

respectively: N, T - S, B, C, and Q, with the low order byte sent first.